## Slide 1: interim NBM sim results
24-06-13

## Slide 2: contents
updates
safety indices
uncertainties in code
next steps

## Slide 3: updates
resolved “heating timeline bug”:
integration of Imasonic transducers into configs
  (chose 100 mm RoC for deeper geometric focus)
initial sim results with Imasonic on NBM in one sbj

## Slide 4: Isppa pattern CTX500 vs. Imasonic
CTX500
Imasonic
(TODO:
bowls should be differently sized)
free water intensity: | 62.2421 W/cm^2 | 62.2421 W/cm^2
max Isppa: | 1.3104 W/cm^2 | 2.0302 W/cm^2
80s
45% DC
5 Hz PRF
(see matlab Figure 1)
(see matlab Figure 2)

## Slide 5: heating timelines
Imasonic
T
CEM43°C
free water intensity: | 62.2421 W/cm^2

## Slide 6: safety indices
Imasonic
free water intensity: | 62.2421 W/cm^2
max Isppa | max_Isppa skull | max_Isppa brain | max_pressure skull | max_pressure brain | max_MI skull | max_MI brain | isppa at_target | avg_isppa around_target | maxT | maxT brain | maxT skull
13.43416 | 3.797333 | 1.303077 | 627219 | 205291.3 | 1.14514 | 0.374809 | 0.7191563 | 0.3278622 | 37.99302 | 37.32696 | 37.94737
8.619675 | 4.125885 | 2.079949 | 653790.2 | 259365.4 | 1.193652 | 0.4735342 | 1.416319 | 0.5046679 | 38.0135 | 37.48237 | 37.98834
according to ITRUSST consensus
(currently still running:
deeper steering AND
free water I: 225 W/cm²
eta: 1 hour)

## Slide 7: Imasonic position variations
Imasonic
(TODO:
bowl should be larger)
free water intensity: | 62.2421 W/cm^2 | 62.2421 W/cm^2
max Isppa: | 1.3104 W/cm^2 | 2.0302 W/cm^2

## Slide 8: uncertainties in code
“expected_focal_distance_mm”?
  transducer-target-dist? value in profiles?
calculation pressure to intensity (from source_amp) correct?
does center PCD need modeling? - NO
correct profile chosen?
“source_phase_deg” in configs?

## Slide 9: next steps
sims of control region (shallow steering) !!!
  currently running in HPC (eta 5 hours)
resolve uncertainties in code
optimize spatial intensity pattern for NBM via target-transducer coords adjustments
  requires few small script adjustments for quick iteration cycles
sims of more subjects for control region variability estimates
apply probabilistic NBM mask for more useful visualization
NBM
control

## Slide 10: updated next steps
free water intensity confirmation
  adapt source_amp for at least 30 W/m²
ask Matthias for: (3 things)
  expected_focal_distance: computed on the fly?
  correct profile chosen?
  source_phase_deg ok to set to all 0?
targeting
  iterate until mostly correct
near-field only – optimization attempt

## Slide 11: problems/bottlenecks
communication – meetings with Martin in future
roadmap to regaining trust in community?